## Slide 1
ANALYTE | SAMPLE | N | MEAN | WITHIN % CV / Target %CV | BETWEEN % CV / Target %CV
Na (mmol/L) | QC L1 QC L2 QC L3 | 25 25 25 | 121 144 175 | 0.88/0.41 0.79/0.44 1.11/0.87 | 1.08/0.86 0.84/0.87 1.11/2.0
K (mmol/L) | QC L1 QC L2 QC L3 | 25 25 25 | 2.97 4.02 6.32 | 1.51/1.73 0.93/0.01 0.98/1.04 | 1.55/1.20 0.93/0.76 1.06/1.10
UREA (mmol/L) | QC L1 QC L2 QC L3 | 25 25 25 | 3.27 15.18 22.55 | 0.97/1.36 0.43/0.80 0.56/0.95 | 2.23/2.59 1.58/2.21 1.09/1.99
CREATININE (µmol/L) | QC L1 QC L2 QC L3 | 25 25 25 | 58 186 490 | 2.51/1.31 1.61/1.62 1.49/1.76 | 4.77/4.19 3.59/2.75 3.16/2.73
Total Protein (g/L) | QC L1 QC L2 QC L3 | 25 25 25 | 45 58 83 | 1.10/0.87 0.78/0.76 0.74/0.62 | 1.52/2.22 0.84/2.15 0.84/1.85
Albumin (g/L) | QC L1 QC L2 QC L3 | 25 25 25 | 29 38 54 | 0.95/0.58 0.32/0.47 0.69/0.61 | 1.15/1.63 0.68/1.20 1.10/1.2
ALP (IU/L) | QC L1 QC L2 QC L3 | 25 25 25 | 73 190 426 | 1.20/0.67 0.66/0.54 0.68/0.82 | 2.43/3.46 1.85/2.62 0.89/2.73
ALT (IU/L) | QC L1 QC L2 QC L3 | 25 25 25 | 25 103 205 | 1.37/2.11 0.72/1.07 0.58/0.61 | 2.03/4.35 0.74/3.16 1.90/3.32
GGT (IU/L) | QC L1 QC L2 QC L3 | 25 25 25 | 27 69 158 | 1.92/1.57 0.74/0.96 0.63/0.80 | 1.92/4.16 1.00/3.07 0.89/3.19
Total Bilirubin (µmol/L) | QC L1 QC L2 QC L3 | 25 25 25 | 15 50 100 | 1.42/1.33 0.95/2.39 1.32/1.25 | 2.05/3.88 1.30/2.54 2.81/2.00
Supplemental table 1
Within and between batch imprecision using IQC material.

## Slide 2
ANALYTE | SAMPLE | N | MEAN | WITHIN % CV / Target %CV | BETWEEN % CV / Target %CV
Albumin (g/L) | QC L1 QC L2 QC L3 | 25 25 25 | 29 38 54 | 0.47/0.58 0.29/0.47 0.69/0.61 | 0.941/1.63 0.94/1.20 1.10/1.20
Calcium (mmol/L) | QC L1 QC L2 QC L3 | 25 25 25 | 1.69 2.58 3.48 | 0.53/0.80 0.64/0.44 0.62/1.34 | 1.18/1.45 1.42/1.44 0.99/1.39
ALP (IU/L) | QC L1 QC L2 QC L3 | 25 25 25 | 73 190 426 | 1.04/0.67 0.60/0.54 0.72/0.82 | 2.33/3.46 2.32/2.62 1.41/2.73
Phosphate (mmol/L) | QC L1 QC L2 QC L3 | 25 25 25 | 0.74 1.43 2.23 | 1.33/0.96 0.65/0.65 0.46/0.61 | 2.58/3.29 1.38/2.30 1.12/1.74
Total Cholesterol (mmol/L) | QC L1 QC L2 QC L3 | 25 25 25 | 2.77 4.33 7.53 | 0.78/0.75 0.83/0.98 0.94/0.74 | 0.92/1.80 1.02/1.88 1.78/1.57
HDL (mmol/L) | QC L1 QC L2 QC L3 | 25 25 25 | 0.65 1.10 1.92 | 0.89/0.74 0.61/0.73 1.07/0.45 | 2.37/1.78 2.11/2.04 1.78/2.15
Triglycerides (mmol/L) | QC L1 QC L2 QC L3 | 25 25 25 | 0.77 1.67 2.65 | 0.78/0.97 1.04/0.67 1.05/0.73 | 1.72/2.26 1.94/1.28 1.83/1.84
B12 (ng/L) | QC L1 QC L2 QC L3 | 25 25 25 | 109 347 571 | 10.13/14.62 3.27/8.65 4.04/8.80 | 10.97/15.10 4.38/9.29 6.91/16.02
FOL (µg/L) | QC L1 QC L2 QC L3 | 25 25 25 | 3.72 6.93 15.71 | 5.30/7.31 4.18/7.17 3.06/7.29 | 8.05/8.22 6.23/10.93 7.50/6.93
FER (µg/L) | QC L1 QC L2 QC L3 | 25 25 25 | 16.8 107.9 176.0 | 2.85/4.70 2.81/5.55 2.03/6.38 | 3.21/7.22 2.81/8.97 5.09/10.31
Supplemental table 1 continued.

## Slide 3
ANALYTE | SAMPLE | N | MEAN | WITHIN % CV / Target %CV | BETWEEN % CV / Target %CV
TSH (mU/L) | QC L1 QC L2 QC L3 | 25 25 25 | 0.04 2.91 17.12 | 2.48/8.17 2.87/4.20 3.95/4.79 | 2.48/8.10 2.87/8.05 3.96/6.65
FT4 (pmol/L) | QC L1 QC L2 QC L3 | 25 25 25 | 5.54 17.10 31.34 | 12.92/11.82 6.53/7.05 2.60/6.54 | 12.92/11.65 6.53/7.34 4.51/5.32
FT3 (pmol/L) | QC L1 QC L2 QC L3 | 25 25 25 | 2.96 5.34 13.82 | 6.29/9.94 4.11/5.79 3.47/6.27 | 8.10/15.03 5.91/10.90 4.75/9.37
Prolactin (mU/L) | QC L1 QC L2 QC L3 | 25 25 25 | 51 144 464 | 4.78/6.82 4.39/4.89 6.90/6.02 | 7.72/10.00 6.54/13.56 6.90/11.04
Cortisol (nmol/L) | QC L1 QC L2 QC L3 | 25 25 25 | 137 404 905 | 5.79/7.48 4.13/6.24 2.80/5.97 | 7.56/5.78 6.10/5.29 6.41/5.61
Supplemental Table 1 continued

## Slide 4
ANALYTE | SAMPLE | N | MEAN | WITHIN % CV |  | BETWEEN % CV | 
 |  |  |  | Patient pool | Manufacturer | Patient pool | Manufacturer
Na (mmol/L) | POOL 1 POOL 2 | 25 25 | 137 150 | 0.60 0.59 | 0.49 | 0.80 0.88 | 0.64
K (mmol/L) | POOL 1 POOL 2 | 25 25 | 4.2 4.6 | 0.67 1.07 | 0.65 | 0.69 1.28 | 0.76
Urea (mmol/L) | POOL 1 POOL 2 | 25 25 | 5.7 15.8 | 1.64 1.51 | 2.28 | 2.38 2.31 | 3.25
Creatinine (µmol/L) | POOL 1 POOL 2 | 25 25 | 65 142 | 2.47 1.60 | 1.12 | 3.01 2.75 | 1.67
Albumin (g/L) | POOL 1 POOL 2 | 25 25 | 37 39 | 0.70 0.74 | 0.58 | 1.11 1.03 | 0.99
Calcium (mmol/L) | POOL 1 POOL 2 | 25 25 | 2.32 2.40 | 0.60 0.80 | 0.65 | 0.89 1.15 | 0.96
ALP (IU/L) | POOL 1 POOL 2 | 25 25 | 110 80 | 1.07 1.31 | 2.1 | 2.25 2.17 | 4.2
Phosphate (mmol/L) | POOL 1 POOL 2 | 25 25 | 1.18 1.22 | 1.09 0.54 | 1.11 | 2.02 2.43 | 1.36
Total Protein (g/L) | POOL 1 POOL 2 | 25 25 | 67 66 | 0.68 0.79 | 1.50 | 1.16 1.16 | 1.80
ALP (IU/L) | POOL 1 POOL 2 | 25 25 | 110 80 | 1.25 1.31 | 2.1 | 2.5 2.17 | 4.2
ALT (IU/L) | POOL 1 POOL 2 | 25 25 | 14 10 | 3.81 4.94 | 2.1 | 4.42 8.2 | 2.7
GGT (IU/L) | POOL 1 POOL 2 | 25 25 | 43 31 | 0.93 1.60 | 1.63 | 1.97 2.66 | 2.4
Total Bilirubin (µmol/L) | POOL 1 POOL 2 | 25 25 | 8.3 5.7 | 2.34 2.72 | 1.03 | 2.97 3.55 | 2.59
Supplemental Table 2
Within and between batch imprecision using pooled patient serum with corresponding imprecision values stated by the manufacturer

## Slide 5
ANALYTE | SAMPLE | N | MEAN | WITHIN % CV |  | BETWEEN % CV | 
 |  |  |  | Patient pool | Manufacturer | Patient pool | Manufacturer
Total Cholesterol (mmol/L) | POOL 1 POOL 2 | 25 25 | 4.54 4.67 | 1.00 0.52 | 0.50 | 1.62 1.64 | 0.80
HDL (mmol/L) | POOL 1 POOL 2 | 25 25 | 1.38 1.22 | 1.47 1.21 | 0.85 | 1.85 2.25 | 1.92
Triglycerides (mmol/L) | POOL 1 POOL 2 | 25 25 | 1.30 1.65 | 0.88 0.97 | 0.70 | 1.25 1.46 | 1.10
Coritsol (nmol/L) | POOL 1 POOL 2 POOL 3 | 25 25 25 | 150 310 674 | 4.76 4.23 2.76 | 6.70 | 7.36 4.23 2.97 | 7.90
PROLACTIN (mU/L) | POOL 1 POOL 2 POOL 3 | 25 25 25 | 101 260 738 | 2.96 3.24 3.36 | 1.61 | 2.96 3.24 4.45 | 6.92
TSH (mU/L) | POOL 1 POOL 2 POOL 3 | 25 25 25 | 0.142 1.617 22.86 | 3.09 2.69 11.18 | 3.0 | 3.67 3.36 11.18 | 2.0
FT4 (pmol/L) | POOL 1 POOL 2 POOL 3 | 25 25 25 | 8.32 11.10 19.14 | 6.13 5.35 2.94 | 4.4 | 6.87 5.35 4.09 | 8.08
B12 (ng/L) | POOL 1 POOL 2 POOL 3 | 25 25 25 | 153 330 745 | 4.87 6.29 9.64 | 5.0 | 5.76 8.89 10.69 | 8.5
FOLATE (µg/L) | POOL 1 POOL 2 POOL 3 | 25 25 25 | 2.43 9.57 63.35 | 8.77 3.42 3.10 | 3.6 | 10.72 6.90 9.66 | 4.7
FERRITIN (µg/L) | POOL 1 POOL 2 POOL 3 | 25 25 25 | 68 135 463 | 5.10 3.36 3.46 | 1.7 | 5.41 3.70 3.46 | 2.5
Supplemental Table 2 continued

## Slide 6
Chart: 

Identity | 130 | 144
 | 139 | 139 | 138 | 136 | 141 | 132 | 139 | 134 | 139 | 142 | 139 | 139 | 139 | 139 | 130 | 142 | 138 | 137 | 138 | 142 | 137 | 131 | 141 | 138 | 136 | 132 | 138 | 142 | 143 | 142 | 139 | 143
Passing-Bablok fit (y = -0.5 + 1 x) | 130.5 | 143.5
Sodium (mmol/L)
Chart: 

Identity | 3.6 | 5.8
 | 4.7 | 4.9 | 4.8 | 5.1 | 5.1 | 5.4 | 5.2 | 4.4 | 4.6 | 5.2 | 5.1 | 4.9 | 4.8 | 4.4 | 4.5 | 4.9 | 4.8 | 4.3 | 4.4 | 4.7 | 5.7 | 4.5 | 4.6 | 5 | 5.4 | 5.1 | 4.6 | 4.6 | 5.3 | 5 | 4.2
Passing-Bablok fit (y = 0.5 + 1 x) | 4.1 | 5.4
Chart: 

Identity | 2 | 11
 | 6.9 | 2.7 | 6.6 | 3.9 | 5.9 | 6.4 | 8.5 | 6.8 | 6.6 | 4.9 | 6.5 | 8.4 | 5 | 10.2 | 4.6 | 6.4 | 6.4 | 6.2 | 6.2 | 8.1 | 7.1 | 2.8 | 4 | 3.6 | 5.2 | 8.7 | 7.1 | 6 | 4.2 | 10.4 | 5.1 | 4.1
Passing-Bablok fit (y = -0.02571 + 1.029 x) | 2.5457 | 10.26
Urea (mmol/L)
Chart: 

Identity | 50 | 175
 | 146 | 53 | 62 | 57 | 56 | 80 | 145 | 106 | 56 | 61 | 72 | 101 | 64 | 109 | 51 | 81 | 62 | 75 | 56 | 54 | 125 | 75 | 62 | 62 | 56 | 93 | 69 | 80 | 67 | 174 | 82 | 63
Passing-Bablok fit (y = -2.658 + 1.012 x) | 48.9467 | 171.3811
Creatinine (µmol/L)
Supplemental Figure 1. Passing Bablok plots showing correlation between venous blood results (x axis) vs TAPII sample results (y axis) for U&E profiles. Regression equations are shown within each plot. N = 32 for each analyte except potassium (N = 31) where 1 out of 32 TAPII samples was flagged as haemolysed by automated LIH indices.

## Slide 7
Supplemental Figure 2. Passing Bablok plots showing correlation between venous blood results (x axis) vs TAPII sample results (y axis) for LFT profiles. Units of concentration on x and y axis are IU/L for ALT and ALP, g/L for total protein and albumin and µmol/L for total bilirubin. Regression equations are shown within each plot. N = 26 for ALP, total protein; N = 25 for Albumin.
N = 24 for ALT and total bilirubin.
Chart: 

Identity | 0 | 300
 | 14 | 15 | 24 | 23 | 23 | 13 | 279 | 116 | 13 | 20 | 57 | 15 | 16 | 31 | 17 | 37 | 23 | 14 | 17 | 20 | 26 | 16 | 29 | 28
Passing-Bablok fit (y = 0.4651 + 0.9767 x) | 11.2093 | 279.814
ALT (IU/L)
Chart: 

Identity | 25 | 250
 | 70 | 55 | 52 | 106 | 60 | 78 | 45 | 237 | 139 | 47 | 178 | 94 | 98 | 73 | 30 | 58 | 75 | 74 | 45 | 119 | 87 | 61 | 67 | 96 | 94 | 115
Passing-Bablok fit (y = -0.4063 + 0.9788 x) | 29.9353 | 240.3696
ALP (IU/L)
Chart: 

Identity | 58 | 80
 | 64 | 64 | 61 | 70 | 78 | 71 | 67 | 58 | 76 | 71 | 78 | 70 | 73 | 67 | 65 | 67 | 62 | 65 | 64 | 69 | 67 | 66 | 63 | 66 | 65 | 65
Passing-Bablok fit (y = -1 + 1 x) | 58 | 78
Total Protein (g/L)
Albumin (g/L)
Total Bilirubin (µmol/L)
Chart: 

Identity | 0 | 200
 | 13 | 16 | 10 | 11 | 13 | 175 | 8 | 6 | 17 | 13 | 26 | 17 | 7 | 14 | 9 | 12 | 13 | 11 | 9 | 11 | 8 | 7 | 9 | 12
Passing-Bablok fit (y = 0 + 1 x) | 6 | 181
Chart: 

Identity | 34 | 46
 | 42 | 39 | 39 | 41 | 40 | 43 | 40 | 34 | 44 | 39 | 46 | 44 | 45 | 39 | 39 | 42 | 40 | 42 | 43 | 44 | 42 | 42 | 41 | 42 | 42
Passing-Bablok fit (y = -1 + 1 x) | 34 | 45

## Slide 8
Supplemental Figure 3. Passing Bablok plots showing correlation between venous blood results (x axis) vs TAPII sample results (y axis) for Bone profiles. Albumin was not included here as all patient requests for bone profile also had a liver function test profile requested, thus ALP and Albumin were only measured within the LFT request. Regression equations are shown within each plot. N = 19 for phosphate. N = 20 for adjusted calcium.
Adjusted Calcium (mmol/L)
Chart: 

Identity | 2.15 | 2.6
 | 2.33 | 2.55 | 2.42 | 2.5 | 2.45 | 2.28 | 2.32 | 2.18 | 2.32 | 2.29 | 2.37 | 2.44 | 2.32 | 2.38 | 2.28 | 2.37 | 2.25 | 2.44 | 2.41 | 2.38
Passing-Bablok fit (y = -0.03 + 1 x) | 2.14 | 2.53
Chart: 

Identity | 0.6 | 1.5
 | 1.21 | 1 | 1.18 | 0.78 | 1.15 | 0.92 | 0.7 | 0.93 | 1.02 | 1.06 | 0.86 | 1.23 | 0.86 | 0.89 | 1.01 | 0.81 | 1.43 | 1.1 | 0.8
Passing-Bablok fit (y = 0.02 + 1 x) | 0.69 | 1.37
Phosphate (mmol/L)

## Slide 9
Supplemental Figure 4. Passing Bablok plots showing correlation between venous blood results (x axis) vs TAPII sample results (y axis) for Lipid profiles. Units of concentration on x and y axis are mmol/L. Regression equations are shown within each plot. N = 21 for HDL and Cholesterol N = 20 for Triglycerides and LDL (Friedwald equation).
Chart: 

Identity | 1 | 5.5
 | 2.8582 | 3.1964 | 1.4364 | 2.4109 | 3.1755 | 4.7918 | 4.6191 | 5.3555 | 2.3173 | 1.7327 | 3.8564 | 2.5127 | 3.4809 | 4.0345 | 2.1927 | 2.0782 | 3.2345 | 3.0955 | 4.0909 | 2.8955
Passing-Bablok fit (y = -0.03336 + 0.9895 x) | 1.4582 | 5.0106
Chart: 

Identity | 0.75 | 4
 | 3.7 | 0.91 | 1.02 | 1.01 | 1.77 | 1.47 | 1.96 | 2.32 | 0.93 | 1.05 | 0.8 | 1.16 | 1.12 | 1.75 | 1.27 | 0.84 | 0.98 | 1.55 | 1.23 | 1.33
Passing-Bablok fit (y = -0.02 + 1 x) | 0.8 | 3.75
Chart: 

Identity | 1 | 3.2
 | 1.56 | 1.69 | 1.87 | 1.28 | 2.04 | 1.82 | 3.04 | 1.51 | 1.73 | 1.29 | 1.47 | 1.91 | 1.61 | 1.49 | 1.4 | 1.7 | 1.36 | 1.51 | 1.27 | 1.65 | 1.03
Passing-Bablok fit (y = -0.02378 + 0.9894 x) | 1.0547 | 3.0928
Chart: 

Identity | 3 | 9
 | 6.1 | 5.48 | 3.18 | 4.91 | 5.8 | 8.5 | 7.02 | 8.14 | 4.03 | 3.68 | 6.13 | 4.65 | 5.48 | 6.23 | 4.45 | 4.47 | 3.82 | 5.19 | 5.07 | 6.3 | 4.53
Passing-Bablok fit (y = -0.07922 + 0.9887 x) | 3.1439 | 8.5519
Total Cholesterol (mmol/L)
TRIGLYCERIDES (mmol/L)
HDL
(mmol/L)
LDL
(mmol/L)

## Slide 10
Supplemental Figure 5. Passing Bablok plots showing correlation between venous blood results (x axis) vs TAPII sample results (y axis) for Thyroid profiles. Units of concentration on x and y axis are mU/L for TSH and pmol/L for FT4. Regression equations are shown within each plot. N = 22 for TSH and N = 21 for FT4.
Chart: 

Identity | 7 | 20
 | 11.2 | 16.1 | 12 | 12.4 | 7.7 | 13.4 | 15.2 | 10.7 | 12.2 | 12.2 | 11.8 | 19.1 | 9.9 | 11.5 | 14.5 | 13.1 | 10.7 | 12.5 | 11.1 | 10.2 | 10.6
Passing-Bablok fit (y = -0.5665 + 1.051 x) | 7.8418 | 19.193
Chart: 

Identity | 0.5 | 4.5
 | 1.72 | 3.748 | 0.507 | 1.353 | 1.097 | 1.779 | 1.28 | 0.577 | 3.206 | 0.971 | 4.163 | 0.645 | 3.322 | 1.015 | 1.506 | 1.192 | 1.526 | 2.312 | 1.54 | 3.766 | 2.864 | 1.626
Passing-Bablok fit (y = 0.02218 + 0.9699 x) | 0.55755 | 4.3468
TSH
(mU/L)

## Slide 11
Supplemental Figure 6. Passing Bablok plots showing correlation between venous blood results (x axis) vs TAPII sample results (y axis) for Cortisol and Prolactin. Units of concentration on x and y axis are nmol/L for cortisol and mU/L for Prolactin and mmol/mol for HbA1c. Regression equations are shown within each plot. N = 19 for Cortisol and Prolactin N= 20 for HbA1c .
Chart: 

Identity | 75 | 375
 | 331 | 317 | 208 | 133 | 179 | 148 | 175 | 219 | 154 | 138 | 286 | 319 | 283 | 158 | 234 | 136 | 175 | 213 | 80
Passing-Bablok fit (y = -8 + 1 x) | 69 | 346
Cortisol
(nmol/L)
Chart: 

Identity | 75 | 350
 | 144 | 135 | 159 | 221 | 105 | 130 | 136 | 123 | 85 | 270 | 204 | 102 | 288 | 265 | 132 | 111 | 95 | 132 | 300
Passing-Bablok fit (y = 5.387 + 0.9964 x) | 90.0766 | 344.146
Chart: 

Identity | 30 | 46
 | 35 | 38 | 40 | 33 | 35 | 32 | 35 | 35 | 40 | 38 | 34 | 40 | 44 | 43 | 31 | 39 | 36 | 35 | 44 | 40
Passing-Bablok fit (y = -7.5 + 1.2 x) | 30.9 | 45.3
HbA1c
(mmol/mol)

## Slide 12
Supplemental Figure 7. Passing Bablok plots showing correlation between venous blood results (x axis) vs TAPII sample results (y axis) for Haematinic profiles. Units of concentration on x and y axis are µg/L for Folate and Ferritin and ng/L for B12. Regression equation and Pearson correlation coefficient are shown within each plot. N = 17 for B12, N = 18 for Folate and N = 15 for Ferritin.
Chart: 

Identity | 150 | 500
 | 306 | 458 | 273 | 451 | 232 | 278 | 348 | 215 | 283 | 165 | 387 | 207 | 232 | 301 | 299 | 255 | 313
Passing-Bablok fit (y = 6.651 + 0.9508 x) | 165.4413 | 461.1522
Chart: 

Identity | 150 | 500
 | 306 | 458 | 273 | 451 | 232 | 278 | 348 | 215 | 283 | 165 | 387 | 207 | 232 | 301 | 299 | 255 | 313
Passing-Bablok fit (y = 6.651 + 0.9508 x) | 165.4413 | 461.1522
B12
(ng/L)
Chart: 

Identity | 2 | 24
 | 6.8 | 17.8 | 14.6 | 6.7 | 8.1 | 6.7 | 23.5 | 6.3 | 10.1 | 20.89 | 14.5 | 13.2 | 7 | 8.1 | 7.8 | 16.4 | 6.5 | 4.2
Passing-Bablok fit (y = 0.03671 + 1.051 x) | 4.0291 | 20.5241
Chart: 

Identity | 2 | 24
 | 6.8 | 17.8 | 14.6 | 6.7 | 8.1 | 6.7 | 23.5 | 6.3 | 10.1 | 20.89 | 14.5 | 13.2 | 7 | 8.1 | 7.8 | 16.4 | 6.5 | 4.2
Passing-Bablok fit (y = 0.03671 + 1.051 x) | 4.0291 | 20.5241
Folate
(
Chart: 

Identity | 0 | 250
 | 24 | 153 | 134 | 114 | 177 | 23 | 202 | 61 | 38 | 241 | 41 | 62 | 234 | 14 | 69
Passing-Bablok fit (y = 3.333 + 0.9611 x) | 15.8278 | 242.65
Chart: 

Identity | 0 | 250
 | 24 | 153 | 134 | 114 | 177 | 23 | 202 | 61 | 38 | 241 | 41 | 62 | 234 | 14 | 69
Passing-Bablok fit (y = 3.333 + 0.9611 x) | 15.8278 | 242.65